ROTOPOL set-up   (4/25/2013)

In case of trouble: email Werner Kaminsky: Kaminsky@chem.washington.edu

Device built for Paul Bollyky <pbollyky@benaroyaresearch.org>

and Patrick Secor <psecor@uw.edu>

source of software: http://128.95.152.162/RotopolHome/RotpolInstall.MSI

Password (requested ones on second start of program)  Rotopol#2011

Motor settings:
	Acceleration:	20
	Speed:		60
	Power:		95%
	Steps/360  deg.	16000
	counterclockwize (if camera’s settings mirror
 			feature is changed, this may 
			need to be actualized)

sin(d) correction: with supplied near perfect circular polarizer, all colors:
1 / 0
1 / 90

Camera: 
	driver:		MD800E
	white:		R:100, G:130, B:130 (avoid blue too high)
	Exposure:	NO autoexposure
			target 120, time 1500, gain 40
	Color:		Gamma 1.0
			Saturation 128
			Contrast 23  (small changes affect 0-offset, 
					default was 22)
	Miscell.		Flip Vertically checked
			60 Hz
			low Frame Rate

Resolution:
	works with 8 or 16 bits settings, use 16 bits if effects are small
	RGB colors checked
	dark level 0.01 (increase if noise level too high or readings irrational 
			on low intensity pixels.)

Connection of device: 	black USB (camera) on right closer to screen than
			the white USB (motor).