Crystal Structure Report for X. XXX

by Werner Kaminsky

Sample Submitted by XXX XXX

ID – XXX

X-ray Crystallography Laboratory

University of Washington

31. July 2026

Data Collection  started:

Documentation finished:

Data backup accessibly through http:

URL: http://cad4.cpac.washington.edu/structures

W.Kaminsky: tel. (206) 543 7585

(x-ray lab):     tel. (206) 543 0210

Email: kaminsky@chem.washington.edu

Understanding crystallographic information is not trivial. If you discuss a structure in a publication, please let the crystallographers proof-read the final draft of that manuscript before you submit it.

For many structures it is true that crystallography is an intellectual performance (like synthesis or elucidation of mechanisms) and not a routine job. In these cases, particularly when the structures contribute significantly to the content of the publication, the crystallographers should be co-authors. Do not think this is not necessary because you pay for structures.

A xxx xxx, measuring 0.xx x 0.xx x 0.xx mm3 was mounted on a loop with oil.  Data was collected at -173oC on a Bruker APEX II single crystal X-ray diffractometer, Mo-radiation.

Crystal-to-detector distance was xx mm and exposure time was xx seconds per frame for all sets.  The scan width was x.xo.   Data collection was xx.x% complete to 25o  in .  A total of  xxx reflections were collected covering the indices, h = -xx to xx, k = -xx to xx,  l = -xx to xx.  xxx reflections were symmetry independent and the Rint = 0.xxxx indicated that the data was of xxx than average quality (0.07).  Indexing and unit cell refinement indicated a xxx xxx lattice.  The space group was found to be x xx  (No.xx).

The data was integrated and scaled using SAINT, SADABS within the APEX2 software package by Bruker.

The data appeared twinned via 180 degrees about [-1 0 1] (CELL_NOW) and after  multi-domain integration (SAINT, SADABS within the APEX2 software package by Bruker) the data was merged utilizing twinabs.

The contribution of disordered solvent to the diffraction pattern was removed with SQUEEZE

Solution by direct methods (SHELXS, SIR97) produced a complete heavy atom phasing model consistent with the proposed structure. The structure was completed by difference Fourier synthesis with SHELXL97., Scattering factors are from Waasmair and Kirfel. Hydrogen atoms were placed in geometrically idealised positions and constrained to ride on their parent atoms with C---H distances in the range 0.95-1.00 Angstrom. Isotropic thermal parameters Ueq were fixed such that they were 1.2Ueq of their parent atom Ueq for CH's and 1.5Ueq of their parent atom Ueq in case of methyl groups. All non-hydrogen atoms were refined anisotropically by full-matrix least-squares.

<Describe structure here>

The structure is of high quality and ready for publication. Table 1 summarizes the data collection details. Figure 1 shows an ORTEP of the asymmetric unit.

<put gif here>

Figure 1. ORTEP of the structure with thermal ellipsoids at the 50% probability level. Disorder and alcohol solvent omitted for clarity.

Table 1: Crystallographic data for the structures provided.

<put table 1 here>

Bruker (2007) APEX2 (Version 2.1-4), SAINT (version 7.34A), SADABS (version 2007/4), BrukerAXS Inc, Madison, Wisconsin, USA.

Sheldrick, G. M. (2005). CELL_NOW. University of Goettingen, Germany.

Bruker (2007) APEX2 (Version 2.1-4), SAINT (version 7.34A), SADABS (version 2007/4), BrukerAXS Inc, Madison, Wisconsin, USA.

Sheldrick, G. M. (2007). TWINABS. University of Goettingen, Germany.

(a) A. Spek, J. Appl. Cryst. (2003), 36, 7-13;

(b) P. van der Sluis & A. L. Spek, Acta Cryst. (1990). A46, 194-201.

(c) Spek, A. L. (2009). Acta. Cryst. D65, 148-155.

(a) Altomare A, Burla C, Camalli M, Cascarano G L, Giacovazzo C, Guagliardi A,   Moliterni AGG, Polidori G, Spagna R. (1999) SIR97: a new tool for crystal structure determination and refinement Journal of Applied Crystallography, 32, 115-119.

(b) Altomare A, Cascarano G L, Giacovazzo C, Guagliardi A. (1993) Completion and refinement of crystal structures with SIR 92. Journal of Applied Crystallography, 26, 343-350.

(a) Sheldrick GM. (1997) SHELXL-97, Program for the Refinement of Crystal Structures. University of Göttingen, Germany.

(b) Sheldrick GM. (2013) Crystal structure refinement with SHELXL. Acta Cryst. (2015). C71, 3-8.

Mackay, S.; Edwards, C.; Henderson, A.; Gilmore, C.; Stewart, N.; Shankland, K.; Donald, A. (1997) MaXus: a computer program for the solution and refinement of crystal structures from diffraction data. University of Glasgow, Scotland,.

Waasmaier, D.; Kirfel, A. (1995) New Analytical Scattering Factor Functions for Free Atoms and Ions. Acta Crystallographica A., 51, 416-430.

Farrugia LJ. (1997) Ortep-3 for Windows. Journal of Applied Crystallography, 30, 565