Crystal Structure Report for K. Goldberg

by Betsy Flowers and Werner Kaminsky

Sample Submitted by Ash Wreight

ID – AW3_57

X-ray Crystallography Laboratory

University of Washington

31. July 2026

Data Collection  started: 12/30/2014

Documentation finished: 1/8/2015

Data backup accessibly through http:

URL: 	http://cad4.cpac.washington.edu/structures

Username: 	xray-user

Password: 	hkl-data

W.Kaminsky: tel. (206) 543 7585

(x-ray lab):     tel. (206) 543 0210

Email: kaminsky@chem.washington.edu

Understanding crystallographic information is not trivial. If you discuss a structure in a publication, please let the crystallographers proof-read the final draft of that manuscript before you submit it.

For many structures it is true that crystallography is an intellectual performance (like synthesis or elucidation of mechanisms) and not a routine job. In these cases, particularly when the structures contribute significantly to the content of the publication, the crystallographers should be co-authors. Do not think this is not necessary because you pay for structures. Even I  have to pay for my own structures for my research.

A yellow piece, cut from a bigger crystal, measuring 0.30 x 0.26 x 0.24 mm3 was mounted on a loop with oil.  Data was collected at -153oC on a Bruker APEX II single crystal X-ray diffractometer, Mo-radiation.

Crystal-to-detector distance was 40 mm and exposure time was 2 seconds per frame for all sets.  The scan width was 0.5o.   Data collection was 92.8% complete to 25o  in .  A total of  430385 reflections were collected covering the indices, h = -22 to 22, k = -22 to 22,  l = -22 to 22.  5798 reflections were symmetry independent and the Rint = 0.0557 indicated that the data was of good (average quality 0.07).  Indexing and unit cell refinement indicated a primitive orthorhombic lattice.  The space group was found to be P n m a (No.62).

Some small amount of twinning, possibly an unavoidable result of cutting the sample from a larger crystal, was treated the usual way (CELL_NOW, twinabs), with small data improvement. The data was integrated and scaled using SAINT within the APEX2 software package by Bruker.

Solution by direct methods (SHELXS, SIR97) produced a complete heavy atom phasing model consistent with the proposed structure. The structure was completed by difference Fourier synthesis with SHELXL97., Scattering factors are from Waasmair and Kirfel. Hydrogen atoms were placed in geometrically idealised positions and constrained to ride on their parent atoms with C---H distances in the range 0.95-1.00 Angstrom. Isotropic thermal parameters Ueq were fixed such that they were 1.2Ueq of their parent atom Ueq for CH's and 1.5Ueq of their parent atom Ueq in case of methyl groups. All non-hydrogen atoms were refined anisotropically by full-matrix least-squares.

Several CH2Cl2 solvents with indication of being multiply disordered were first included in the refinement, but a better treatment of the situation required their contribution to the diffraction pattern better being removed via SQUEEZE. The molecule contains a mirror plane through N3, C23, C22, Ir, both Br and C1, C4, C5.

The structure is of high quality and ready for publication. Table 1 summarizes the data collection details. Figure 1 shows an ORTEP of the asymmetric unit.

Figure 1. ORTEP of the structure with thermal ellipsoids at the 50% probability level.

Table 1: Crystallographic data for the structures provided.

Empirical formula 	C41 H35 Br2 F3 Ir N5

Formula weight 	1006.76

Temperature 	120(2) K

Wavelength 	0.71073 Å

Crystal system 	Orthorhombic

Space group 	P n m a

Unit cell dimensions	a = 16.4831(12) Å	= 90°.

b = 16.8327(13) Å	= 90°.

c = 16.9552(13) Å	 = 90°.

Volume	4704.3(6) Å3

Z	4

Density (calculated)	1.421 Mg/m3

Absorption coefficient	4.578 mm-1

F(000)	1960

Crystal size	0.30 x 0.26 x 0.24 mm3

Theta range for data collection	2.11 to 28.50°.

Index ranges	-22<=h<=22, -22<=k<=22, -22<=l<=22

Reflections collected	168741

Independent reflections	5798 [R(int) = 0.0829]

Completeness to theta = 25.00°	94.1 %

Max. and min. transmission	0.4062 and 0.3404

Refinement method	Full-matrix least-squares on F2

Data / restraints / parameters	5798 / 0 / 254

Goodness-of-fit on F2	1.052

Final R indices [I>2sigma(I)]	R1 = 0.0652, wR2 = 0.1453

R indices (all data)	R1 = 0.0943, wR2 = 0.1542

Largest diff. peak and hole	2.761 and -3.349 e.Å-3

Sheldrick, G. M. (2005). CELL_NOW. University of Goettingen, Germany.

Sheldrick, G. M. (2007). TWINABS. University of Goettingen, Germany.

Bruker (2007) APEX2 (Version 2.1-4), SAINT (version 7.34A), SADABS (version 2007/4), BrukerAXS Inc, Madison, Wisconsin, USA.

(a) Altomare A, Burla C, Camalli M, Cascarano G L, Giacovazzo C, Guagliardi A,   Moliterni AGG, Polidori G, Spagna R. (1999) SIR97: a new tool for crystal structure determination and refinement Journal of Applied Crystallography, 32, 115-119.

(b) Altomare A, Cascarano G L, Giacovazzo C, Guagliardi A. (1993) Completion and refinement of crystal structures with SIR 92. Journal of Applied Crystallography, 26, 343-350.

Sheldrick GM. (1997) SHELXL-97, Program for the Refinement of Crystal Structures. University of Göttingen, Germany.

Mackay, S.; Edwards, C.; Henderson, A.; Gilmore, C.; Stewart, N.; Shankland, K.; Donald, A. (1997) MaXus: a computer program for the solution and refinement of crystal structures from diffraction data. University of Glasgow, Scotland,.

Waasmaier, D.; Kirfel, A. (1995) New Analytical Scattering Factor Functions for Free Atoms and Ions. Acta Crystallographica A., 51, 416-430.

A. Spek, J. Appl. Cryst. (2003), 36, 7-13; P. van der Sluis & A. L. Spek, Acta Cryst. (1990). A46, 194-201.

Farrugia LJ. (1997) Ortep-3 for Windows. Journal of Applied Crystallography, 30, 565