Crystal Structure Report for K Goldberg

by Sarah Flowers and Werner Kaminsky

Sample Submitted by Karena Smoll

ID – KAS_1_90

X-ray Crystallography Laboratory

University of Washington

28. July 2026

Data Collection  started: 10/22/2014

Documentation finished: 10/27/2014

Data backup accessibly through http:

URL: http://cad4.cpac.washington.edu/structures

W.Kaminsky: tel. (206) 543 7585

(x-ray lab):     tel. (206) 543 0210

Email: kaminsky@chem.washington.edu

Understanding crystallographic information is not trivial. If you discuss a structure in a publication, please let the crystallographers proof-read the final draft of that manuscript before you submit it.

For many structures it is true that crystallography is an intellectual performance (like synthesis or elucidation of mechanisms) and not a routine job. In these cases, particularly when the structures contribute significantly to the content of the publication, the crystallographers should be co-authors. Do not think this is not necessary because you pay for structures. Even I  have to pay for my own structures for my research.

A colorless chard, measuring 0.06 x 0.06 x 0.05 mm3 was mounted on a loop with oil.  Data was collected at -23oC on a Bruker APEX II single crystal X-ray diffractometer, Mo-radiation.

Crystal-to-detector distance was 40 mm and exposure time was 60 seconds per frame for all sets.  The scan width was 0.5o.   Data collection was 99.9% complete to 25o  in .  A total of  xxx reflections were collected covering the indices, h = -14 to 14, k = -18 to 18,  l = -19 to 18.  35896 reflections were symmetry independent and the Rint = 0.1401 indicated that the data came from a very small weakly diffracting sample.  Indexing and unit cell refinement indicated a primitive monoclinic lattice.  The space group was found to be P 21/n  (No. 14).

The data was integrated and scaled using SAINT, SADABS within the APEX2 software package by Bruker.

Solution by direct methods (SHELXS, SIR97) produced a complete heavy atom phasing model consistent with the proposed structure. The structure was completed by difference Fourier synthesis with SHELXL97., Scattering factors are from Waasmair and Kirfel. Hydrogen atoms were placed in geometrically idealised positions and constrained to ride on their parent atoms with C---H distances in the range 0.95-1.00 Angstrom. Isotropic thermal parameters Ueq were fixed such that they were 1.2Ueq of their parent atom Ueq for CH's and 1.5Ueq of their parent atom Ueq in case of methyl groups. All non-hydrogen atoms were refined anisotropically by full-matrix least-squares.

The structure contains voids of 108 Å3.  Such voids are only stable in crystals with a network of hydrogen bonds. Indeed, a dimeric hydrogen system exists between O3 –H3O – O2’ - C25’ – O3’ – H3O’ O2. Lowering the temperature seems to destabilize the network and thus, the sample exhibits a structural phase transition of first order between 225 and 250 K. The transition seed energy is high because the transition occurs with a time lack of 5 or more hours.

The structure is of high quality and ready for publication. Table 1 summarizes the data collection details. Figure 1 shows an ORTEP of the asymmetric unit.

Figure 1. ORTEP of the structure with thermal ellipsoids at the 50% probability level.

Table 1: Crystallographic data for the structures provided.

Empirical formula 	C25 H44 O3 P2 Pd

Formula weight 	560.94

Temperature 	250(2) K

Wavelength 	0.71073 Å

Crystal system 	Monoclinic

Space group 	P 21/n

Unit cell dimensions	a = 12.284(4) Å	= 90°.

b = 15.339(5) Å	= 105.671(16)°.

c = 16.105(5) Å	 = 90°.

Volume	2921.8(16) Å3

Z	4

Density (calculated)	1.275 Mg/m3

Absorption coefficient	0.766 mm-1

F(000)	1176

Crystal size	0.06 x 0.06 x 0.05 mm3

Theta range for data collection	2.17 to 25.62°.

Index ranges	-14<=h<=14, -18<=k<=18, -19<=l<=18

Reflections collected	35896

Independent reflections	5461 [R(int) = 0.1401]

Completeness to theta = 25.00°	99.9 %

Max. and min. transmission	0.9627 and 0.9555

Refinement method	Full-matrix least-squares on F2

Data / restraints / parameters	5461 / 0 / 293

Goodness-of-fit on F2	0.985

Final R indices [I>2sigma(I)]	R1 = 0.0512, wR2 = 0.0858

R indices (all data)	R1 = 0.1273, wR2 = 0.1075

Largest diff. peak and hole	0.464 and -0.640 e.Å-3

Bruker (2007) APEX2 (Version 2.1-4), SAINT (version 7.34A), SADABS (version 2007/4), BrukerAXS Inc, Madison, Wisconsin, USA.

(a) Altomare A, Burla C, Camalli M, Cascarano G L, Giacovazzo C, Guagliardi A,   Moliterni AGG, Polidori G, Spagna R. (1999) SIR97: a new tool for crystal structure determination and refinement Journal of Applied Crystallography, 32, 115-119.

(b) Altomare A, Cascarano G L, Giacovazzo C, Guagliardi A. (1993) Completion and refinement of crystal structures with SIR 92. Journal of Applied Crystallography, 26, 343-350.

Sheldrick GM. (1997) SHELXL-97, Program for the Refinement of Crystal Structures. University of Göttingen, Germany.

Mackay, S.; Edwards, C.; Henderson, A.; Gilmore, C.; Stewart, N.; Shankland, K.; Donald, A. (1997) MaXus: a computer program for the solution and refinement of crystal structures from diffraction data. University of Glasgow, Scotland,.

Waasmaier, D.; Kirfel, A. (1995) New Analytical Scattering Factor Functions for Free Atoms and Ions. Acta Crystallographica A., 51, 416-430.

Farrugia LJ. (1997) Ortep-3 for Windows. Journal of Applied Crystallography, 30, 565