Crystal Structure Report for J Kovacs

by Maike Blakely & Werner Kaminsky

Sample Submitted by Maike Blakely

ID – MB1050_2

X-ray Crystallography Laboratory

University of Washington

25. March 2016

Data Collection  started: 03/22/2016

Documentation finished: 03/25/2016

Data backup accessibly through http:

URL: http://cad4.cpac.washington.edu/structures

W.Kaminsky: tel. (206) 543 7585

(x-ray lab):     tel. (206) 543 0210

Email: kaminsky@chem.washington.edu

Understanding crystallographic information is not trivial. If you discuss a structure in a publication, please let the crystallographers proof-read the final draft of that manuscript before you submit it.

For many structures it is true that crystallography is an intellectual performance (like synthesis or elucidation of mechanisms) and not a routine job. In these cases, particularly when the structures contribute significantly to the content of the publication, the crystallographers should be co-authors. Do not think this is not necessary because you pay for structures. Even I  have to pay for my own structures for my research.

A colorless prism, measuring 0.40 x 0.30 x 0.10 mm3 was mounted on a glass capillary with oil.  Data was collected at -173oC on a Bruker APEX II single crystal X-ray diffractometer, Mo-radiation.

Crystal-to-detector distance was 40 mm and exposure time was 10 seconds per degree for all sets.  The scan width was 0.5o.   Data collection was 99.8% complete to 25o  in (.  A total of  412633 reflections were collected covering the indices, -26<=h<=26, -34<=k<=34, -20<=l<=20.  18971 reflections were symmetry independent and the Rint = 0.1118 indicated that the data was appropriate.  Indexing and unit cell refinement indicated a primitive orthorhombic lattice.  The space group was found to be P n a 21  (No.33).

The data was integrated and scaled using SAINT, SADABS within the APEX2 software package by Bruker.

Solution by direct methods (SHELXS, SIR97) produced a complete heavy atom phasing model consistent with the proposed structure. The structure was completed by difference Fourier synthesis with SHELXL97., Scattering factors are from Waasmair and Kirfel. Hydrogen atoms were placed in geometrically idealised positions and constrained to ride on their parent atoms with C---H distances in the range 0.95-1.00 Angstrom. Isotropic thermal parameters Ueq were fixed such that they were 1.2Ueq of their parent atom Ueq for CH's and 1.5Ueq of their parent atom Ueq in case of methyl groups. All non-hydrogen atoms were refined anisotropically by full-matrix least-squares.

The structure is of high quality and ready for publication. Table 1 summarizes the data collection details. Figure 1 shows an ORTEP of the asymmetric unit.

Figure 1. ORTEP of the structure with thermal ellipsoids at the 50% probability level.

Table 1: Crystallographic data for the structures provided.

Empirical formula 	C90 H75 I Na O5 P5

Formula weight 	1541.24

Temperature 	100(2) K

Wavelength 	0.71073 Å

Crystal system 	Orthorhombic

Space group 	P n a 21

Unit cell dimensions	a = 19.6017(13) Å	= 90°.

b = 25.4669(17) Å	= 90°.

c = 15.1649(10) Å	 = 90°.

Volume	7570.2(9) Å3

Z	4

Density (calculated)	1.352 Mg/m3

Absorption coefficient	0.589 mm-1

F(000)	3176

Crystal size	0.40 x 0.30 x 0.10 mm3

Theta range for data collection	1.31 to 28.46°.

Index ranges	-26<=h<=26, -34<=k<=34, -20<=l<=20

Reflections collected	412633

Independent reflections	18971 [R(int) = 0.1118]

Completeness to theta = 25.00°	99.8 %

Max. and min. transmission	0.9435 and 0.7986

Refinement method	Full-matrix least-squares on F2

Data / restraints / parameters	18971 / 1 / 919

Goodness-of-fit on F2	1.041

Final R indices [I>2sigma(I)]	R1 = 0.0551, wR2 = 0.1540

R indices (all data)	R1 = 0.0642, wR2 = 0.1622

Absolute structure parameter	-0.044(14)

Largest diff. peak and hole	1.690 and -2.244 e.Å-3

# Endnotes

Bruker (2007) APEX2 (Version 2.1-4), SAINT (version 7.34A), SADABS

(version 2007/4), BrukerAXS Inc, Madison, Wisconsin, USA.

(a) Altomare, A.; Burla, C.; Camalli M.; Cascarano L.; Giacovazzo C.; Guagliardi A.;   Moliterni A.G.G.; Polidori G.; Spagna R. J. Appl. Cryst. 1999, 32, 115-119; (b) Altomare, A., Cascarano, G., Giacovazzo, C., Guagliardi, A., J. Appl. Cryst. 1993, 26, 343.

Sheldrick, G. M. SHELXL-97: Program for the Refinement of Crystal Structures 1997 University of Gottingen, Germany.

Mackay, S.; Edwards, C.; Henderson, A.; Gilmore, C.; Stewart, N.; Shankland, K.; Donald, A.; MaXus: a computer program for the solution and refinement of crystal structures from diffraction data. University of Glasgow, Scotland, 1997.

Waasmaier, D.; Kirfel, A. Acta Crystallogr. A. 1995, 51, 416.

Farrugia, L. J. Appl. Cryst. 1997, 30, 565