Crystal Structure Report for J Kovacs

by Maike Blakely & Werner Kaminsky

Sample Submitted by Maike Blakely

ID – MB1085_1b

X-ray Crystallography Laboratory

University of Washington

02. June 2016

Data Collection  started: 5/23/2016

Documentation finished: 6/2/2016

Data backup accessibly through http:

URL: http://cad4.cpac.washington.edu/structures

W.Kaminsky: tel. (206) 543 7585

(x-ray lab):     tel. (206) 543 0210

Email: kaminsky@chem.washington.edu

Understanding crystallographic information is not trivial. If you discuss a structure in a publication, please let the crystallographers proof-read the final draft of that manuscript before you submit it.

For many structures it is true that crystallography is an intellectual performance (like synthesis or elucidation of mechanisms) and not a routine job. In these cases, particularly when the structures contribute significantly to the content of the publication, the crystallographers should be co-authors. Do not think this is not necessary because you pay for structures.

An orange prism, measuring 0.07 x 0.01 x 0.01 mm3 was mounted on a loop with oil.  Data was collected at -173oC on a Bruker APEX II single crystal X-ray diffractometer, Mo-radiation.

Crystal-to-detector distance was 40 mm and exposure time was 240 seconds per frame for all sets.  The scan width was 1o.   Data collection was 98.8 % complete to 25o  in (.  A total of  24836 reflections were collected covering the indices, -12<=h<=12, -12<=k<=12, -17<=l<=17.  3061 reflections were symmetry independent and the Rint = 0.0942 indicated that the data was of o.k..  Indexing and unit cell refinement indicated a primitive orthorhombic lattice.  The space group was found to be P 21 n m  (No. 31).

The data was integrated and scaled using SAINT, SADABS within the APEX2 software package by Bruker.

Solution by direct methods (SHELXS, SIR97) produced a complete heavy atom phasing model consistent with the proposed structure. The structure was completed by difference Fourier synthesis with SHELXL97., Scattering factors are from Waasmair and Kirfel. Hydrogen atoms were placed in geometrically idealised positions and constrained to ride on their parent atoms with C---H distances in the range 0.95-1.00 Angstrom. Isotropic thermal parameters Ueq were fixed such that they were 1.2Ueq of their parent atom Ueq for CH's and 1.5Ueq of their parent atom Ueq in case of methyl groups. All non-hydrogen atoms were refined anisotropically by full-matrix least-squares.

The structure appears pseudo-tetragonal. However, trying basic symmetries like P 4 or P(4 lead to non-sensical models. The structure seems to be twinned (0.00 -1.00  0.00 -1.00  0.00  0.00  0.00  0.00 -1.00). In addition, strong disorder appears to raise the apparent symmetry. The Chlorides are disordered with a small percentage of iodide.

At this stage, no higher symmetry was found.

The structure is of moderate quality and may not be considered for publication. Table 1 summarizes the data collection details. Figure 1 shows an ORTEP of the asymmetric unit.

Figure 1. ORTEP of the structure with thermal ellipsoids at the 50% probability level. Disorder omitted for clarity.

Table 1: Crystallographic data for the structures provided.

Empirical formula 	C20 H36 Cl1.78 Fe3 I0.22 O4 S4

Formula weight 	727.30

Temperature 	100(2) K

Wavelength 	0.71073 Å

Crystal system 	Orthorhombic

Space group 	P 21 n m

Unit cell dimensions	a = 10.2163(13) Å	= 90°.

b = 10.1992(12) Å	= 90°.

c = 14.7735(18) Å	 = 90°.

Volume	1539.4(3) Å3

Z	2

Density (calculated)	1.569 Mg/m3

Absorption coefficient	2.064 mm-1

F(000)	744

Crystal size	0.07 x 0.01 x 0.01 mm3

Theta range for data collection	1.38 to 25.38°.

Index ranges	-12<=h<=12, -12<=k<=12, -17<=l<=17

Reflections collected	24836

Independent reflections	3061 [R(int) = 0.0942]

Completeness to theta = 25.00°	98.6 %

Max. and min. transmission	0.9797 and 0.8690

Refinement method	Full-matrix least-squares on F2

Data / restraints / parameters	3061 / 242 / 220

Goodness-of-fit on F2	1.091

Final R indices [I>2sigma(I)]	R1 = 0.0849, wR2 = 0.1822

R indices (all data)	R1 = 0.1090, wR2 = 0.1944

Absolute structure parameter	0.36(9)

Largest diff. peak and hole	1.322 and -0.804 e.Å-3

# Endnotes

Bruker (2007) APEX2 (Version 2.1-4), SAINT (version 7.34A), SADABS (version 2007/4), BrukerAXS Inc, Madison, Wisconsin, USA.

(a) Altomare A, Burla C, Camalli M, Cascarano G L, Giacovazzo C, Guagliardi A,   Moliterni AGG, Polidori G, Spagna R. (1999) SIR97: a new tool for crystal structure determination and refinement Journal of Applied Crystallography, 32, 115-119.

(b) Altomare A, Cascarano G L, Giacovazzo C, Guagliardi A. (1993) Completion and refinement of crystal structures with SIR 92. Journal of Applied Crystallography, 26, 343-350.

(a) Sheldrick GM. (1997) SHELXL-97, Program for the Refinement of Crystal Structures. University of Göttingen, Germany.

(b) Sheldrick GM. (2013) Crystal structure refinement with SHELXL. Acta Cryst. (2015). C71, 3-8.

Mackay, S.; Edwards, C.; Henderson, A.; Gilmore, C.; Stewart, N.; Shankland, K.; Donald, A. (1997) MaXus: a computer program for the solution and refinement of crystal structures from diffraction data. University of Glasgow, Scotland,.

Waasmaier, D.; Kirfel, A. (1995) New Analytical Scattering Factor Functions for Free Atoms and Ions. Acta Crystallographica A., 51, 416-430.

Farrugia LJ. (1997) Ortep-3 for Windows. Journal of Applied Crystallography, 30, 565