Crystal Structure Report for K. Goldberg

by Maike Blakey & Werner Kaminsky

Sample Submitted by Zoha Soyed

ID – ZS_3_25

X-ray Crystallography Laboratory

University of Washington

31. July 2026

Data Collection  started: 8/30/2017

Documentation finished: 9/19/2017

Data backup accessibly through http:

URL: http://cad4.cpac.washington.edu/structures

W.Kaminsky: tel. (206) 543 7585

(x-ray lab):     tel. (206) 543 0210

Email: kaminsky@chem.washington.edu

Understanding crystallographic information is not trivial. If you discuss a structure in a publication, please let the crystallographers proof-read the final draft of that manuscript before you submit it.

For many structures it is true that crystallography is an intellectual performance (like synthesis or elucidation of mechanisms) and not a routine job. In these cases, particularly when the structures contribute significantly to the content of the publication, the crystallographers should be co-authors. Do not think this is not necessary because you pay for structures.

A yellow prism, measuring 0.39 x 0.12 x 0.12 mm3 was mounted on a loop with oil.  Data was collected at -173oC on a Bruker APEX II single crystal X-ray diffractometer, Mo-radiation.

Crystal-to-detector distance was 40 mm and exposure time was 10 seconds per frame for all sets.  The scan width was 0.5o.   Data collection was 100% complete to 25o in .  A total of 107980 reflections were collected covering the indices, -35<=h<=35, -17<=k<=17, -13<=l<=13.  4210 reflections were symmetry independent and the elevated Rint = 0.0265 indicated that the data was of better than average quality (0.07).   Indexing and unit cell refinement indicated a primitive monoclinic lattice.  The space group was found to be C 2/c (No. 15).

The data was integrated and scaled using SAINT, SADABS within the APEX2 software package by Bruker.

Solution by direct methods (SHELXS, SIR97) produced a complete heavy atom phasing model consistent with the proposed structure. The structure was completed by difference Fourier synthesis with SHELXL97., Scattering factors are from Waasmair and Kirfel. Hydrogen atoms were placed in geometrically idealised positions and constrained to ride on their parent atoms with C---H distances in the range 0.95-1.00 Angstrom. Isotropic thermal parameters Ueq were fixed such that they were 1.2Ueq of their parent atom Ueq for CH's and 1.5Ueq of their parent atom Ueq in case of methyl groups. All non-hydrogen atoms were refined anisotropically by full-matrix least-squares.

It was discovered by chance that the material can crystallize also in a triclinic polymorph with three molecules per asymmetric unit. It is possible that some residual polymorphism could have coexisted in the selected sample causing some problems with rest electron density, however, by limiting the maximum scattering angle to 2 = 50 degrees, most of the problems were taken care of and the presented result is ready for publication.

Table 1 summarizes the data collection details. Figure 1 shows an ORTEP of the asymmetric unit.

Figure 1. ORTEP of the structure with thermal ellipsoids at the 50% probability level.  Solvent omitted.

Table 1: Crystallographic data for the structures provided.

Empirical formula 	C41 H58 Cl4 Ir2 N4 O10

Formula weight 	1293.11

Temperature 	100(2) K

Wavelength 	0.71073 Å

Crystal system 	Monoclinic

Space group 	C 2/c

Unit cell dimensions	a = 29.768(2) Å	= 90°.

b = 14.8097(11) Å	= 104.976(4)°.

c = 11.1859(8) Å	 = 90°.

Volume	4763.8(6) Å3

Z	4

Density (calculated)	1.803 Mg/m3

Absorption coefficient	5.863 mm-1

F(000)	2536

Crystal size	0.390 x 0.120 x 0.120 mm3

Theta range for data collection	1.416 to 25.345°.

Index ranges	-35<=h<=35, -17<=k<=17, -13<=l<=13

Reflections collected	8251

Independent reflections	4210 [R(int) = 0.0265]

Completeness to theta = 25.000°	96.3 %

Refinement method	Full-matrix least-squares on F2

Data / restraints / parameters	4210 / 3 / 289

Goodness-of-fit on F2	1.187

Final R indices [I>2sigma(I)]	R1 = 0.0385, wR2 = 0.0880

R indices (all data)	R1 = 0.0668, wR2 = 0.1090

Largest diff. peak and hole	2.893 and -1.296 e.Å-3

Bruker (2007) APEX2 (Version 2.1-4), SAINT (version 7.34A), SADABS (version 2007/4), BrukerAXS Inc, Madison, Wisconsin, USA.

(a) Altomare A, Burla C, Camalli M, Cascarano G L, Giacovazzo C, Guagliardi A,   Moliterni AGG, Polidori G, Spagna R. (1999) SIR97: a new tool for crystal structure determination and refinement Journal of Applied Crystallography, 32, 115-119.

(b) Altomare A, Cascarano G L, Giacovazzo C, Guagliardi A. (1993) Completion and refinement of crystal structures with SIR 92. Journal of Applied Crystallography, 26, 343-350.

(a) Sheldrick GM. (1997) SHELXL-97, Program for the Refinement of Crystal Structures. University of Göttingen, Germany.

(b) Sheldrick GM. (2013) Crystal structure refinement with SHELXL. Acta Cryst. (2015). C71, 3-8.

Mackay, S.; Edwards, C.; Henderson, A.; Gilmore, C.; Stewart, N.; Shankland, K.; Donald, A. (1997) MaXus: a computer program for the solution and refinement of crystal structures from diffraction data. University of Glasgow, Scotland,.

Waasmaier, D.; Kirfel, A. (1995) New Analytical Scattering Factor Functions for Free Atoms and Ions. Acta Crystallographica A., 51, 416-430.

Farrugia LJ. (1997) Ortep-3 for Windows. Journal of Applied Crystallography, 30, 565